## Slide 1: MIRA36 at Lindenberg Experiences and results of ten years of operation
Ulrich Görsdorf and many co-authors
Meteorologisches Observatorium Lindenberg, DWD

## Slide 2
Content
System characteristics
System performance (reliability, accuracy,..)
Calibration issues
Applications

## Slide 3
Motivation
Providing long term data sets of cloud parameters, which cannot be delivered by conventional observations, as basis for
Continuous evaluation of NWP models
Climatology studies

## Slide 4
TX: Magnetron, PEP 30 kW, f = 35.1 GHz
Cassegrain antenna with polarization filter,
vertically pointed, d = 1.9 m, G = 54.5 dBi
Digital receiver FN = 6.2 dB
Pulse width 200 ns
Pulse repetition freq. 5 kHz
N. of spectra averaging 200 (40)
Averaging time 10 s (2 s spectra)
Range 0.15 – 15 km
FFT length 256
Sensitivity: -55 dBZ
Stored data: Doppler moments (mmclx), Compressed spectra (zspc, zspca), Full spectra (pds, ones a week)
MIRA36 at Lindenberg

## Slide 5
Commisioning in Nov. 2003
First commercial system by METEK
Two (!) magnetron transmitter
1 m Cassegrain antenna
Water cooling for magnetron
Transmitter I
Transmitter I
Waveguide switch
Receiver
Power supply

## Slide 6
November 2003
Delivery of the radar by Metek

## Slide 7
System upgrades and
modifications
November 2003 | Commissioning, start of test operation
1. April 2004 | Start of routine operation
7./8. July 2004 | Installation clutter fence
April 2006 | Modification of Magnetron coolant from water to air
November 2006 | Data processing by mmclx
February 2010 | System upgrade (digital receiver, 1.9m antenna)
February-May 2012 | Change of carrier frequency
September 2013 | Installation of clutter fence

## Slide 8
Fast Fourier
Transformation
Doppler Spectra
*.pds
Noise estimation, Clutter filtering
Calculation of Doppler Moments
Multi-Peak-Analysis
(spc2dmp)
*.dmp
Calculation of Ze
Spektrum-based
Target-Classification
(mmclx)
Target-Classification
2. Step (cloudnet)
Flags for
ice, cloud droplets,
rain/drizzle,
insects, aerosol
*.mmclx
Retrieval of cloud
parameters
(cloudnet)
liquid water content
ice water content
cloud fraction
*.nc
Calculation of
Doppler Moments
*.zspc
Compression
of spectra
*.pdd
*.nc
Plankton
Filtering
Calculation
of cloud boundaries
*.dat
DSP
Post-Processing
10 sec
2 sec
10 sec
30 sec
30 sec
2 sec
10 sec
2 sec
Additional Data
Temp., humidity profile
Backscatter profile Ceiometer
LWP Radiometer
rain gauge
Ceilometer cloud base
Signal-Processing
Data processing
www.cloud-net.org

## Slide 9: Instrumentation with potential for Cloudnet
Radiosonde, Vaisala RS92
4 times a day
GRUAN Lead Centre
GSN –surface station local observer (24/7)
Ka-band Radar, MIRA36 (24/7)
Raman Lidar RAMSES,
355 nm, 24/on demand
482 MHz Wind Profiler Radar (24/7)
Microwave Profiler (24/7)
Ceilometer CHM15k, LD40 (24/7)
BSRN (24/7)
Whole Sky Imager (24/7),
Nubiscope (24/7)
Doppler Lidar HALO(24/7)
Ceilometer CL31, CL51, LD40 (24/7)

## Slide 10
Content
System characteristics
System performance (reliability, accuracy,..)
Calibration issues
Applications

## Slide 11
Availability
Frequency adjustment
Receiver upgrade, 1.9 m antenna
Defective power supply
Defective thermistor, compressor
83.1
94.6
98.2
91.8
96.7
98.8
83.7
97.3
61.2
99.5
MTBF(2004-2013): 41 days MTBF(2010-2013): 102 days (250 days hardware)

## Slide 12
Frequency distribution of
moments and LDR, 2013

## Slide 13
Time series of monthly means
Reflectivity
ΔZ10yr = 0.24 dB
ΔZ10yr = 1.08 dB
ΔZ10yr = -0.12 dB
ΔZ10yr = -2.88 dB

## Slide 14
Time series of monthly means
Reflectivity

## Slide 15
Time series of monthly means
Spectral width

## Slide 16
Time series of monthly means
LDR

## Slide 17
Unwanted/ suspicious signals
Clouds with unatural high
LDR, chaff ?
Ground clutter
without clutter fence,
Moving tractor on farmland
Suspicious signals above
cloud top , sidelobe effect,
spillover fence
Snow on the antenna, antenna
heating unintentionally turned off

## Slide 18
Content
System characteristics
System performance (reliability, accuracy,..)
Calibration issues
Applications

## Slide 19: Background
good accuracy of reflectivity and LDR is essential for a precise estimation of microphysical parameters

## Slide 20
„Budget“ - Calibration

## Slide 21
Max. ΔZe = 3.1 dB
Prob. ΔZe = 1.3 dB
→ Need for independent verification of calibration

## Slide 22
Using MRR as reference (?)

## Slide 23: Rain
simple experiment to study the influence of rain on Z and LDR
day with a preferably homogenous cloud and a significant LDR

## Slide 24: Results
the effect of rain depends on the strenght of wetting
Dish and plastic cap of Cassegrain reflector are wetted:
- Z drops down by 5 to 15 dBz
- LDR jump up by 3 to 5 dB
Only dish is wetted:
- Z drops down by 3 dBz
- no changes of LDR

## Slide 25
Content
System characteristics
System performance (reliability, accuracy,..)
Calibration issues
Applications

## Slide 26: Mean monthly cloud cover

## Slide 27
low
middle
high
Mean monthly cloud cover
f(cloud heigth)

## Slide 28: Histogram of cloud fraction for each model level
Relative Probability, %
> 20
Model
Measurement
2008 - 2010

## Slide 29: Parametrisation of stratiform clouds
Large discontinuity in cloud fraction distribution above 5 km
Large discontinuity in cloud fraction distribution above 5 km
Large discontinuity in cloud fraction distribution above 5 km
Mean cloud fraction (red lines) and their probability distribution function (in %) for COSMODE
forecasts (left) and radar observations (middle). The cloud fraction differences (in %) are shown on
the right.
Correction of cloud cover Nstrat for subvisible ice clouds above p = 500 hPa
a, b: empirical parameters (a=0.2, b=5 x 10-5), qi: ice water content (iwc)
qi: ice water content (iwc)
a, b: empirical parameter
Routine a=0.2, b=5 x 10-5
Observation a=0.1, b=8 x 10-6

## Slide 30
Relative Probability, %
> 20
Model, improved corr. For subvisible ice clouds
Messung
Model, Routine
Jan. – May 2011
Histogram of cloud fraction for each model level

## Slide 31
Aircraft icing by supercooled droplets

## Slide 32
COSMO
Observation
Underestimation of supercooled droplets by COSMO

## Slide 33: Modification of cloud parametrisation in COSMO
COSMO
Aerosol
Ice & SLW
Ice
Observation
COSMO (old)
COSMO (new)
December 12, 2012 Lindenberg
 | Felix Rieper, Ulrich Görsdorf
 | Towards 3D prediction of supercooled liquid water for aircraft icing: Modifications of the microphysics in COSMO Submitted to Met. Zeitschrift

## Slide 34: Summary
Continuous operation with high availability (MTBF …)
unique cloud parameters for the model validation
Data have been used in several studies
Accuracy of budget calibration has bee aprooved with 1.3 dBz, but a indepented calibration would be very valuable and is still unsolved
Calibration
Software, version control
Data format (netcdf cf 1.6)
Spectra storage, netcdf, compression
Issues, which should be discussed

## Slide 35
Differences of cloud base heigth
Between Radar and Ceilometer

## Slide 36: Detektion unterkühlten Wassers (Methode: Cloudnet – Retrieval)
Radar Z ~ D6
Ceilometer β ~ D2
Mikrowellenradiometer

## Slide 37: Detektion unterkühlten Wassers (Methode: Cloudnet – Retrieval)
Wassertropfen
wenn θ < 0°C, dann
Schicht unterkühlten Wassers
Radiometer → Wassergehalt
Illingworth at al., 2007
Hogan/O’Connor, 2004
Ceilometer
Radarreflektivität, dBz
Rückstreuintensität, b.E.
Radar

## Slide 38: Ziel: Verbesserung Parameterisierung COSMO-EU /ICON
Reduzierung des Bergeron-Findeisen Prozesses
  Reduzierung des Depositionswachstums von Eis und Schnee
  Reduzierung der Zahl der Eiskristalle
Reduzierung der Regengefrierrate
Details siehe Poster
F. Rieper et al. : Improvement of supercooled liquid water prediction
Operationelle Vereisungsprognose durch eigenständiges Ein-Säulen Modell (ADWICE)
© Koninklijk Nederlands Meteorologisch Instituut (KNMI)

## Slide 39
December 12, 2012 Lindenberg
Auswirkung auf SLW-Pfad
COSMO-EU (old)
COSMO-EU (new)
Observation
good retrieval